## Slide 1: Optical Systems Interactions using FDS
Description of each
How they work
Advantages and disadvantages of each
Choosing which to use

## Slide 2: Optical system requirements
Source
Cell
Imaging
Optics
Detector
Radial resolution
Sensitivity
Range
Linearity
Time to scan

## Slide 3: Absorbance Optical System Source overview
Source
Torroidal grating
Wavelength selection
Source intensity

## Slide 4: Synchronizing sources
Rotor pulse
Sample pulse
Reference pulse
1
Magnet
Magnet detector
Clock
Lamp pulse

## Slide 5: Absorbance detector
Radial motion
Photomultiplier
Tube
Top View
Cell

## Slide 6: Absorbance signal processing
PMT
i
Light
Transconductance
amplifier
V
PGA
Vxg
Sample and Hold
and V/f converter
f
Counter
16-bit count
Gating
Data
Acquisition
Board
Rotor timing pulse
hv set
Gain

## Slide 7: Absorbance signal
A = log (Io/I)
  Low Io results in high noise level
  Save data as intensities
Noise is high frequency
  Stochastic- rms 1/N1/2
    Diminishing returns beyond N=9
  Dependent on signal levels
  Best between 0.1-1.0 A

## Slide 8: Absorbance system facts
Xenon lamp intensity ‘spiky’ and blue
  Maximum light intensity 230 nm
  Work at high intensity regime
Maximum repetition rate 100 Hz
  Limits scan rate above 6000 rpm
Scales gain on reference intensity
  Uses 6.5 cm
  ‘Invert’ sample and reference may not work

## Slide 9: When to use Absorbance
Need selectivity
Added sensitivity
Cannot dialyze sample

## Slide 10: Absorbance system ‘gotchas’
Low intensity!
  Weak lamp output
  Buffer absorbs
Schlieren distortion at high gradients
  Light deviated out of system
  Typically overestimate A(r)
Slit to lamp image tracking
  Steep absorbance curve
  Test holmium oxide cell
r
r

## Slide 11: Interference optical system
Copyright Beckman 1996
Lightsource

## Slide 12: Interference source
Collimating
Lens
Wavefronts
Laser Diode
Mask

## Slide 13: Optical path difference
Optical path difference
Copyright Beckman 1996

## Slide 14: Pathlength difference
Mask
Reference
no
Sample
n
In phase
Out of phase

## Slide 15: Interference optical system
Interference
Copyright Beckman 1996

## Slide 16: Pathlength difference leads to fringe displacement
d1
d2
d1’
d2’
Condensing
lens
DY= Dn l /l
d1’-d2’=l/2
d1-d2=0

## Slide 17: Pathlength difference
Mask
Reference
no
Sample
n
In phase
Out of phase
r
d1 – d2

## Slide 18: Optical path difference
Imaging optics
Copyright Beckman 1996

## Slide 19: Superposition of images
Camera lens
Cell
Cylinder lens
Fringes

## Slide 20: Interference optical system facts
Everything that contributes to d1-d2 will show up in fringe displacement
  Only need to alter d1-d2 by l*0.01
Must use sapphire windows
Noise tends to be systematic, low frequency
∆Y= ∆n l /l
  Shorter wavelength laser will increase sensitivity
Radial resolution 10 μm

## Slide 21: Interference system ‘gotchas’
Incomplete dialysis of sample
  Buffer components contribute equally and are in much higher concentrations
  Watch out for detergents!
Improper optical focus and alignment
  Image looks fine… just inaccurate!
Inappropriate blank correction
  Water blanks should be used
  Not needed for ∆c/∆t, Sedanal

## Slide 22: When to use interference optics
Buffer absorbs
Sample does not
Precision required
Extinction coefficient varies
Short columns

## Slide 23: AU-FDS optics Confocal fluorescence optics
PMT
Dichroic mirror
Spatial filter
Fluorescent event
Collimation
Objective lens
Expansion
Laser
Dichroic filter

## Slide 24: AU-FDS
Excitation: 15 mW 488 nm
  Target: fluorescein-like dyes and GFP
    Protein intrinsic fluorescence not accessible
Emission: >505 nm
  Bandpass filters
Optics move as a unit
  Radial positioning by stepping motor
Objective lens focused by stepping motor
PMT
Dichroic Mirror
Spatial filter
Fluorescent event
Wavelength
filter
Objective lens
Laser
Lens assy.
Spatial filter

## Slide 25: AU-FDS Synchronizing
Uses a continuous light source
Acquire light intensities and rotor timing pulse
  Asynchronously with the rotor spinning
  For several turns of the rotor for averaging
Sort out the data in computer memory
Fluorescence
Rapid-scan absorbance
Intensity from sample 1
One rotation
Rotor timing pulse
Intensity

## Slide 26: Data averaging Multiple rotations and multiple intensities
Data for all samples
  Speeds up scanning
Signal averaging
  Multiple intensities/rev
  Multiple revolutions
1
2
3

## Slide 27: Data Collection One radius, several rotations
FDS
One rotation
Radial
steps
FDS
Signal from samples
One rotation

## Slide 28: Applications for FDS
Extrinsically labeled compounds
  488 nm excitation (fluorescein, Alexa, etc.)
  Green fluorescent protein
High selectivity and sensitivity
  Trace materials
  Complex fluids (serum, sputum, urine)
Tight associations (Kd < 10-9 M)
Broad concentration range
  3-4 orders of magnitude

## Slide 29: Fluorescence system gotchas
F = Io Qec
  Q is quantum yield
    Quenching- remove O2
    Photobleaching
Have to label a molecule
  Changes to function?
Sample adsorption at low concentrations
  Meniscus, windows, centerpiece
  Makes F  c difficult to obtain
Inner filter effect
Signal roll-off at base
  Can use FC 43

## Slide 30: GFP in 100% Serum

## Slide 31: Equilibrium Scans of 40 nM GFP in Serum
Notes: Rotor Speed (20k RPM) _ X% serum (diluted with PBS)

## Slide 32: Summary

## Slide 33: Summary comparison
Sensitivity Radial Resolution Scan time When to Use | Absorbance 0.1 OD 20-50 μm 1-5 minutes Selectivity Sensitivity Non-dialyzable | Fluorescence 100 pM fluorescein 20-50 μm 60 seconds (all cells) Selectivity Sensitivity Non-dialyzable Small quantities | Interference 10-6 Δn 10 μm Seconds Buffer absorbs Sample doesn’t Variable ε Accurate C Short column equilibrium

## Slide 34: Optical system choices by sample type and desired initial concentrations
 | Absorbance | Interference | Fluorescence
Sensitivity Range Precision | 0.1 OD 2-3 logs Good | 0.05 mg/ml 3-4 logs Excellent | 100 pM fluorescein 3-4 logs Good
Protein
Choice of optics
1 A230 or 280
1 mg/ml
5 nM fluorescein
Polysaccharide
Interference optics
C ~1 mg/ml
5 nM fluorescein
Nucleic Acid
Absorbance optics
1 A260
5 nM fluorescein

## Slide 35: High concentration systems
Volume fraction and viscosity

## Slide 36: High concentration effects
Molecule k affects sj
Solvent flow, viscosity
Coupled flows
50 microns

## Slide 37: Hydrodynamics: volume fraction Φ
David Goodsell
R
Separated ~20% of R
But fraction of what?
Simple CC f = 40%
SCC max = 52.3%
Face CC f = 56%
FCC max = 74%

## Slide 38: 40 nM GFP in E. coli Lysate Macromolecular protein solvation
Rachel Kroe
s ↑ as M ↑
s ↓ as ρ↑η↑

## Slide 39: Ab:Ab interactions Tracer IgG in high concentration IgG background
IgGs in 10 mM His pH 6.0
IgG A
IgG B
IgG C
poly IgG
s ↑ as M ↑
s ↓ as ρ↑η↑

## Slide 40: IgGs solvate other IgGs
Sue Chase
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Trace Background
mAbB in mAbA Yes
mAbB in mAbB No
mAbB in mAbC No
mAbA in pIgG Yes
Trace Background
mAbC in mAbA Yes
mAbC in mAbB No
mAbC in mAbC No
mAbA in pIgG Yes
Trace Background
mAbA in mAbA Yes
mAbA in mAbB No
mAbA in mAbC No
mAbA in pIgG Yes
A in B or C
A in A
A in pIg
B in B or C
B in A
B in pIg
C in B or C
C in A
C in pIg
A interacts with high A
B,C interact with high A
A,B,C interact with Poly IgG
s ↑ as M ↑
s ↓ as ρ↑η↑

## Slide 41: 440 pM labeled RANKL + varying mAb in PBS Lattice overshoot kinetics
Hill JJ & Laue TM (2015) “Protein assembly in serum and the differences from assembly in buffer.” Methods in Enz. 562, pp. 501-527
s ↑ as M ↑
s ↓ as ρ↑η↑

## Slide 42: 19 nM labeled Anti RANKL + 85 nM RANKL
in PBS
in Serum
IgG binds to serum component
No large complexes!!
Hill JJ & Laue TM (2015) “Protein assembly in serum and the differences from assembly in buffer.” Methods in Enz. 562, pp. 501-527

## Slide 43: Albumin binds fluorescein weakly
100 mg/mL BSA
0.1 mg/mL BSA
PBS, pH 7, 1 μg/mL Fluorescein
Jon Kingsbury

## Slide 44: Fluorescein in PBS and serum

## Slide 45: Example of J-O effect
550 nM GFP in serum